## Slide 1: PLUS Tutorial: Ultrasound probe calibration using fCal
https://plustoolkit.org
Last Updated: 11/12/2018
By Adam Rankin

## Slide 2: Concept
Goal: Determine the Image-to-Probe transform.
Use the correspondence of known wire positions in 3D and segmented wire positions in the US image.
N-wires are attached to a precisely manufactured calibration phantom (see image below) – see phantom building tutorial for instructions for building a phantom

## Slide 3: Coordinate systems

## Slide 4: Hardware setup
Device set used in this tutorial:
SonixTouch with GPS extension
Tracker tools
  Probe
  Reference
  Stylus
fCal 2.x calibration phantom
SonixGPS side cover removed for accessing all connectors

## Slide 5: Hardware setup
If using optical tracking: fixing the markers on the phantom

## Slide 6: Software overview
Download Plus applications releases from https://plustoolkit.github.io/download.html
Start fCal.exe (PlusApp_x.x.x.xxxx/bin/fCal.exe)
Select your Device set as “SonixTouch GPS L14 + Reference + Stylus + Phantom x.x”
Click Connect
Check if all the devices are shown “OK”
Start Stylus calibration
Phantom registration
Temporal Calibration
Spatial calibration

## Slide 7: Software setup
Download Plus applications releases from https://plustoolkit.github.io/download.html
Start fCal.exe (PlusApp_x.x.x.xxxx/bin/fCal.exe)
Select your Device set as “SonixTouch GPS L14 + Reference + Stylus + Phantom x.x”
*L14 refers to .STL file for the probe only but not limits to the probe calibration, the user can use this config file for both C5 and L14

## Slide 8: Software setup
Click Connect and check if all the devices are shown “OK”

## Slide 9: Stylus calibration
Attach an EM sensor to any pointing object.
Click “Start”. fCal collects 200 different points (defined in configuration file). Stylus tip should stay touching a single point while pivoting the stylus so that the sensor traverses a sphere surface
Save the calibration, select the directory to be saved
7
5
6

## Slide 10: Phantom registration
Click “Record point” or pivot the stylus about the current landmark to register 8 predefined landmark points
Model points are defined in the configuration XML file
Tracked points are collected by the stylus. Then save the config file again

## Slide 11: Temporal calibration
Check if the depth is set as desired (10 cm, 12 cm, etc…). Place transducer in a water tank and a preview of the line segmentation should appear. Slowly move the transducer up-down (approximately 2s period).
Continue moving the transducer and then click “Start”. Toolbox automatically collects 100 images for validation and 200 images for calibration (defined in configuration file). Images that cannot segment the line are discarded. See a sample video here: http://www.youtube.com/watch?v=cZq_ZAxA49c
Click “Show Plots” to check result. The two lines should have an offset in the top image, and should be roughly aligned in the bottom image.

## Slide 12: Temporal calibration
If you acquire data using a framegrabber then you have to define the region of interest to exclude text and other information in the image
The same ROI is used for temporal calibration as for spatial calibration, so the ROI can be set in the same place, in the Spatial calibration tab, as shown above

## Slide 13: Temporal calibration
Temporal calibration requires an image where the bottom of the water tank appears as a (not too thick) line
If the line is too thick then adjust the imaging parameters (decrease the power and gain) and/or put a diffuse reflective surface (rubber, cloth, or plastic sheet in the tank)
This “line” is way too thick
Nice, clearly defined line

## Slide 14: Spatial calibration
Place the phantom in the water tank
Align the transducer to have the marked side face to letter “M” on the phantom
Select “Spatial calibration”
Marking on the transducer
Letter “M” marked on the phantom

## Slide 15: Spatial calibration
Select to edit parameters
Select “Region of interest” and pink frame appears on the screen. Move the probe until 9 points of wires appears on the screen.
Click “Freeze Image” then adjust the size of the frame to cover those points. If you do not have much noise in your image, expand the ROI to make it easier to acquire frames.

## Slide 16: Spatial calibration
Set the mass candidates from “80 to 40”
Select “Spacing” to adjust the distance between points (30x20). Align the horizontal alignment tool to align with the top left and top right points. Align the vertical alignment tool to align with the top left and bottom left points. Now the dots should be recognized and turn to green. Move the probe up-down, left-right or until the green dots appear to be stable. Then click “Apply and Close”
Click “Start” and scan the phantom in all degrees of freedom while keeping the points visible. Save the calibration.

## Slide 17: Spatial calibration
Only the fiducials should be visible in the image, the background should be almost black.
Background is OK here.
Background is way too bright here. Decrease the dynamic range, decrease gain, adjust TGCs, increase threshold to make the background darker.
This is way too big. Decrease the gain and dynamic range.
Fiducial line intersection points should appear as small circular-shaped blobs.
This is about OK size, but could be even smaller.
Reduce reflections by slightly changing the transducer angle or putting sound absorbing material in the water container (such as rubber or sponge).
There is a strong reflection from here that brightens up a large section of the image.
Good fiducial size
Keep the transducer's axial (direction of sound travel) axis orthogonal to the wires for optimal reflection from the wires. This dotted line should be approximately horizontal.

## Slide 18: Spatial calibration
The calibration results are shown on the left side.
Select “Show all devices” to view the simulation results
Re-calibrate if the wires are not align properly.
Good Calibration
Bad Calibration