## Slide 1
M
F
N
U
M
F
M
F
M
F
Transducer axes:
M = marked
U = unmarked = -marked
F = far
N = near = -far
marker
marker

## Slide 2
M
F
N
U
Marker symbol (U) is shown at the marked/near corner

## Slide 3
(x=0, y=480)
(x=640, y=0)
(x=640, y=480)
(x=0, y=0)
x direction (towards x increase)
y
(towards y increase)
M
F
N
U
x=U
y=F
=> UltrasoundImageOrientation = UF

## Slide 4
Image size:
640x480 pixels
U
N
F
M
FanAngles[1] = -60 (deg)
FanAngles[2] = 60 (deg)
FanDepth = 450 (pixels)
FanOrigin = 320 40 (pixels)
MF coordinate
System origin
(0, 0)
Fan is always defined in the MF coordinate system!
If the UltrasoundImageOrientation in the stored file is not MF, and the file is loaded into a software that ignores the UltrasoundImageOrientation field (such as ImageJ, Slicer, Paraview) then the XY positions and orientations shown in the software has to be transformed. E.g., if UltrasoundImageOrientation=MN and the FanOrigin appears in the (320,440) position in ImageJ then the FanOrigin in the XML config file shall be (320,40).
Defining the region of interest for volume reconstruction - curvilinear transducer

## Slide 5
Image size: 640x480 pixels
U
N
F
M
ClipRectangleSize (first component) = 600 (pixels)
ClipRectangleOrigin = 20 40 (pixels)
MF coordinate
system origin
(0, 0)
ClipRectangle is always defined in the MF coordinate system!
Clipping rectangle must be always defined. Clipping fan is optional. All pixels that are outside the clipping rectangle or the clipping fan will be ignored in volume reconstruction.
ClipRectangleSize (second component) = 360 (pixels)
Defining the region of interest for volume reconstruction - linear transducer

## Slide 6
F
M
ThetaStartDeg = -60.0
ThetaStopDeg = 60.0
RadiusStopMm = 82.0
(0, 0)
Defining the transducer geometry for RF scan conversion - curvilinear transducer
RadiusStartMm = 10.0
OutputImageStartDephMm = -3.0
OutputImageSizePixel[1]=616
OutputImageSizePixel[0]=820
TransducerGeometry= " CURVILINEAR"
Output image size in mm = OutputImageSizePixel * OutputImageSpacingMmPerPixel

## Slide 7
ImagingDepthMm = 55.0
Defining the transducer geometry for RF scan conversion - linear transducer
TransducerGeometry="LINEAR"
F
M
(0, 0)
OutputImageSizePixel[1]=616
TransducerWidthMm = 38.0
OutputImageSizePixel[0]=820
Output image size in mm = OutputImageSizePixel * OutputImageSpacingMmPerPixel