## Slide 1: Implementation of the PLUS open-source toolkit for translational research of ultrasound-guided intervention systems
Andras Lasso, Tamas Heffter, Csaba Pinter, Tamas Ungi, and Gabor Fichtinger
Laboratory for Percutaneous Surgery (Perk Lab) School of Computing Queen’s University, Kingston, ON, Canada

## Slide 2: Introduction
PLUS – Public software Library for UltraSound imaging
Goal: facilitate rapid prototyping of ultrasound-guided intervention systems for translational clinical research
Acquire, process, transfer synchronized ultrasound image and position data
History: utilizing parts of SynchroGrab library (Boisvert, Gobbi, et al., 2008)
Open-source: BSD license, no strings attached, citation appreciated
Probe
Image
Needle
Reference
Tracker
Phantom model
Ultrasound
computer
3D Slicer
computer

## Slide 3: Data acquisition - tracking
Ascension EM tracker
NDI Aurora, Polaris, and Certus optical and electromagnetic trackers
Claron MicronTracker optical tracker
Brachy steppers (CMS Accuseed, Burdette Medical systems, CIVCO)
PhidgetSpatial inertial measurement device
CHRobotics inertial measurement device
3dConnexion SpaceNavigator 3D mouse
OpenIGTLink (for BrainLab, Siemens MRI scanners, and other compatible devices)
Software devices: file source, US simulator

## Slide 4: Data acquisition - imaging
Ultrasonix: B-mode & RF (through research interface)
BK ProFocus: B-mode & RF (through research interface)
ImagingControl framegrabbers
Epiphan framegrabbers
Video for Windows devices
OpenIGTLink (for MUSiiC, Siemens MRI scanners, and other compatible devices)
Other software devices: file source, US simulator

## Slide 5: Sequence metafile
Extension to the Meta IO standard image file format (*.mha, *.mhd)
Images readable by many existing applications (3D Slicer, ParaView, ImageJ, ITKSnap, …)
Extra information for tracking/reconstruction
  Image orientation
  Transforms from tracker
  Tracker and image status
  Timestamps
  etc.
Used by all applications of Plus
ObjectType = Image
NDims = 3
BinaryData = True
BinaryDataByteOrderMSB = False
CompressedData = True
DimSize = 820 616 97
ElementType = MET_UCHAR
…
UltrasoundImageOrientation = MF
Seq_Frame0000_ProbeToTrackerTransform = 0.956683 -0.263308 0.124204 -190.886 0.269031 0.962616 -0.0315089 -98.0911 -0.111264 0.0635588 0.991756 -1949.07 0 0 0 1
Seq_Frame0000_ProbeToTrackerTransformStatus = OK
Seq_Frame0000_ReferenceToTrackerTransform = 0.260287 0.964665 0.0408926 -208.632 -0.96117 0.262899 -0.0838866 -115.511 -0.0916731 -0.0174702 0.995636 -1992.7 0 0 0 1
Seq_Frame0000_ReferenceToTrackerTransformStatus = OK
Seq_Frame0000_StylusToTrackerTransform = 1 -8.43482e-005 0.000190641 0.370779 8.39344e-005 1 0.000146179 0.316166 -0.00019069 -0.000146138 1 -0.0640528 0 0 0 1
Seq_Frame0000_StylusToTrackerTransformStatus = INVALID
Seq_Frame0000_Timestamp = 345.627957
Seq_Frame0000_ImageStatus = OK
Seq_Frame0001_ProbeToTrackerTransform = 0.956528 -0.263235 0.125549 -190.803 0.268591 0.962861 -0.0275255 -97.8538 -0.113641 0.0600503 0.991706 -1948.58 0 0 0 1
Seq_Frame0001_ProbeToTrackerTransformStatus = OK
…
Seq_Frame0096_StylusToTrackerTransform = 1 0.000362791 -8.01739e-005 -0.106056 -0.000363818 1 -0.000369765 -0.839888 8.00139e-005 0.000370927 1 0.0862409 0 0 0 1
Seq_Frame0096_StylusToTrackerTransformStatus = INVALID
Seq_Frame0096_Timestamp = 355.783014
Seq_Frame0096_ImageStatus = OK
ElementDataFile = LOCAL
xœìÝw...

## Slide 6: Ultrasound image orientation
Specify the orientation of the X and Y axes of the image
Standard orientation in Plus: B-mode => MF, RF-mode => FM

## Slide 7: Temporal calibration
Goal: determine time offset between data streams
Extract the same information (e.g., position) and find the time offset that leads to maximum correlation of the signals
Use accurate clock (1ms is achievable with multimedia timers)
Timestamp filtering: detect delayed data, reduce jitter
Limit maximum speed
without temporal calibration
with temporal calibration
Frame index
time

## Slide 8: Spatial calibration
Goal: determine the PROBE to IMAGE transform

## Slide 9: Spatial calibration
Goal: determine the PROBE to IMAGE transform
Determine STYLUS to STYLUS TIP transform (pivot calibration)
Determine PHANTOM to PHANTOM REFERENCE transform (landmark registration)
Determine the IMAGE to PHANTOM transform (known fiducial line positions in the phantom, segmented visible line positions in the image)
PHANTOM
PHANTOM REFERENCE
STYLUS
STYLUS TIP
PROBE
IMAGE
TRACKER

## Slide 10: Spatial calibration
Tutorials with all data, models, tips&tricks.
Performing tracked ultrasound probe calibration using fCal
How to build an fCal calibration phantom

## Slide 11: Spatial calibration
Generic description of the phantom

## Slide 12: Processing
RF to B-mode conversion: brightness conversion and scan conversion for linear and curvilinear transducers
Volume reconstruction: with hole filling, multi-threaded
Ground truth
Reconstructed volume without hole filling
Reconstructed volume with hole filling

## Slide 13: fCal application
Spatial and temporal calibration
Basic visualization
Recording
Volume reconstruction

## Slide 14: PlusServer application
Console application (no GUI)
Typically runs on the ultrasound computer
Broadcasts data through OpenIGTLink (e.g., to 3D Slicer)

## Slide 15: Performance
Calibration accuracy
  Temporal: standard deviation of repeated lag measurements is a couple of msec
  Spatial: fiducial line reprojection mean error is about 0.5mm (Ultrasonix L14-5 transducer with SonixGPS), total error is about 1mm
Data acquisition: fCal can typically keep up with the data acquisition speed (30-60 fps full-frame, 100-150 fps reduced depth and/or sector size)
OpenIGTLink broadcasting: 10-20 fps
Visualization in 3D Slicer: 10 fps (due to default OpenIGTLinkIF refresh rate)

## Slide 16: Development process
Source control, tickets, releases, messaging (www.assembla.com/spaces/plus)
Standardized build environment (using CMake): automatically download and configure all required software components
Documentation: wiki, doxygen
Automatic tests: CTest, CDash, Sikuli (=> QtTest?)

## Slide 17
PLUS: https://www.assembla.com/spaces/plus
Source code, binary releases, documentations, tutorials, etc.
Slicer IGT: https://www.assembla.com/spaces/SlicerIGT
3D Slicer extensions for additional visualization and analysis.
Contact: Andras Lasso (lasso@cs.queensu.ca) http://perk.cs.queensu.ca
Thank you!

## Slide 18: Appendix

## Slide 19: System overview
hardware
Video interface
Tracker interface
Ultrasound scanner
OpenIGTLink
Navigation system
OpenIGTLink
3D Slicer extension manager
Registration
Segmentation
...
3D Slicer
Temporal and spatial calibration
Visualization
SlicerIGT
Transform recorder
Live ultrasound
PLUS
software
SlicerRT
DICOM-RT import
Dose comparison
VTK, ITK, CTK, QT, DCMTK, …
...
...
ProstateNav
Perk Tutor
MRI scanner
Custom needle guide
Developed in the PerkLab
Developed by collaborators
Custom phantoms

## Slide 20
Source: https://subversion.assembla.com/svn/plus/trunk/doc/overview/PlusOverview.pptx