## Slide 1
Transducer axes:
M = marked
U = unmarked = -marked
F = far
N = near = -far
A = ascending = cross(M,F)
D = descending = -ascending
M
F
N
U
marker
M
F
marker
A
M
F
D
A
M
F
D
For 3D probes only
M
F
F
M

## Slide 2
M
F
N
U
Marker symbol (U) is shown at the marked/near corner

## Slide 3
(x=0, y=480)
(x=640, y=0)
(x=640, y=480)
(x=0, y=0)
x direction (towards x increase)
y
(towards y increase)
M
F
N
U
x=U
y=F
=> UltrasoundImageOrientation = UF

## Slide 4
Image size: 640x480 pixels
U
N
F
M
ClipRectangleSize (first component) = 600 (pixels)
ClipRectangleOrigin = 20 40 (pixels)
MF coordinate
system origin
(0, 0)
ClipRectangleSize (second component) = 360 (pixels)
Defining clipping rectangle for volume reconstruction

## Slide 5
U
N
F
M
FanAnglesDeg[1] = -60 (deg)
FanAnglesDeg[2] = 60 (deg)
FanRadiusStartPixel = 50 (pixels)
FanOrigin = 320 40 (pixels)
MF coordinate
System origin
(0, 0)
Defining clipping fan for volume reconstruction
FanRadiusEndPixel = 450 (pixels)
Image size: 640x480 pixels

## Slide 6
F
M
ThetaStartDeg = -60.0
ThetaStopDeg = 60.0
RadiusStopMm = 82.0
(0, 0)
Defining the transducer geometry for RF scan conversion - curvilinear transducer
RadiusStartMm = 10.0
OutputImageStartDephMm = -3.0
OutputImageSizePixel[1]=616
OutputImageSizePixel[0]=820
TransducerGeometry= " CURVILINEAR"
Output image size in mm = OutputImageSizePixel * OutputImageSpacingMmPerPixel
Obsolete! (only for Plus-2.0.x)
In current versions use TransducerCenterPixel:
TransducerCenterPixel[0] = OutputImageSizePixel[0]/2
TransducerCenterPixel[1] = (RadiusStartMm –OutputImageStartDepthMm) / OutputImageSpacingMmPerPixel[1]

## Slide 7
F
M
ThetaStartDeg = -60.0
ThetaStopDeg = 60.0
RadiusStopMm = 82.0
(0, 0)
Defining the transducer geometry for RF scan conversion - curvilinear transducer
RadiusStartMm = 10.0
Transducer CenterPixel[1] = 53
OutputImageSizePixel[1]=616
OutputImageSizePixel[0]=820
TransducerGeometry= " CURVILINEAR"
Output image size in mm = OutputImageSizePixel * OutputImageSpacingMmPerPixel
Plus 2.1.x and above
TransducerCenterPixel[0] = 410

## Slide 8
ImagingDepthMm = 55.0
Defining the transducer geometry for RF scan conversion - linear transducer
TransducerGeometry="LINEAR"
F
M
(0, 0)
OutputImageSizePixel[1]=616
TransducerWidthMm = 38.0
OutputImageSizePixel[0]=820
Output image size in mm = OutputImageSizePixel * OutputImageSpacingMmPerPixel
Transducer CenterPixel[1] = 15
TransducerCenterPixel[0] = 150