## Slide 1: Implementation of the PLUS open-source toolkit for translational research of ultrasound-guided intervention systems
Andras Lasso, Tamas Heffter, Csaba Pinter, Tamas Ungi, Adam Rankin and Gabor Fichtinger
Laboratory for Percutaneous Surgery (Perk Lab) School of Computing Queen’s University, Kingston, ON, Canada

## Slide 2: Introduction
PLUS – Public software Library for UltraSound imaging
Goal: facilitate rapid prototyping of ultrasound-guided intervention systems for translational clinical research
Acquire, process, transfer synchronized ultrasound image and position data
History: utilizing parts of SynchroGrab library (Boisvert, Gobbi, et al., 2008)
Open-source: BSD license, no strings attached, citation appreciated
Probe
Image
Needle
Reference
Tracker
Phantom model
Ultrasound
computer
3D Slicer
computer

## Slide 3: Data acquisition – pose tracking
Electromagnetic trackers:
  Ascension
  NDI Aurora
Optical trackers:
  NDI Polaris, Certus
  Claron MicronTracker
Surgical navigation systems
  Medtronic StealthStation (through StealthLink)
  BrainLab (through OpenIGTLink)

## Slide 4: Data acquisition – positioning devices
Prostate brachytherapy steppers
  CIVCO EX II
  CMS Accuseed
  Burdette Medical systems
Robotic devices:
  Kuka Lightweight (through OpenIGTLink)
  Intuitive DaVinci (work in progress)

## Slide 5: Data acquisition – ultrasound
Through native research/OEM interface:
  Ultrasonix (including RF data)
  BK (including RF data) through OEM and CameraLink interfaces
  Telemed (including RF data)
  Interson USB probes
  Philips (4D TEE probes)
Any systems with analog or digital video output through framegrabber
Ultrasound B-mode image simulator
MUSiiC toolkit components

## Slide 6: Data acquisition – other imaging
Surgical microscopes, such as Zeiss Pentero (through framegrabber)
Video endoscopes (through framegrabber)
Cameras (through framegrabber)
Webcams, USB cameras (through Microsoft Media Foundation of Video for Windows interface)
Framegrabbers and other imaging devices
  All Epiphan framegrabbers
  All ImagingControl framegrabbers
  Microsoft Media Foundation compatible devices
  Video for Windows compatible devices
Siemens MRI scanners (real-time image acquisition and scan plane control, through OpenIGTLink)

## Slide 7: Data acquisition – other
ThorLabs optical spectrometers
Arduino devices (analog/digital input/output)
Inertial measurement devices:
  PhidgetSpatial
  CHRobotics
3dConnexion SpaceNavigator 3D mouse
OpenIGTLink interface:
  Supported by various hardware and software devices
  C++, Java, Matlab implementations
File sources: for replaying recorded navigation and imaging data (for hardware simulation and testing)

## Slide 8: Sequence metafile
Extension to the Meta IO standard image file format (*.mha, *.mhd)
Images readable by many existing applications (3D Slicer, ParaView, ImageJ, ITKSnap, …)
Extra information for tracking/reconstruction
  Image orientation
  Transforms from tracker
  Tracker and image status
  Timestamps
  etc.
Used by all applications of Plus
Matlab reader/writers are provided
ObjectType = Image
NDims = 3
BinaryData = True
BinaryDataByteOrderMSB = False
CompressedData = True
DimSize = 820 616 97
ElementType = MET_UCHAR
…
UltrasoundImageOrientation = MF
Seq_Frame0000_ProbeToTrackerTransform = 0.956683 -0.263308 0.124204 -190.886 0.269031 0.962616 -0.0315089 -98.0911 -0.111264 0.0635588 0.991756 -1949.07 0 0 0 1
Seq_Frame0000_ProbeToTrackerTransformStatus = OK
Seq_Frame0000_ReferenceToTrackerTransform = 0.260287 0.964665 0.0408926 -208.632 -0.96117 0.262899 -0.0838866 -115.511 -0.0916731 -0.0174702 0.995636 -1992.7 0 0 0 1
Seq_Frame0000_ReferenceToTrackerTransformStatus = OK
Seq_Frame0000_StylusToTrackerTransform = 1 -8.43482e-005 0.000190641 0.370779 8.39344e-005 1 0.000146179 0.316166 -0.00019069 -0.000146138 1 -0.0640528 0 0 0 1
Seq_Frame0000_StylusToTrackerTransformStatus = INVALID
Seq_Frame0000_Timestamp = 345.627957
Seq_Frame0000_ImageStatus = OK
Seq_Frame0001_ProbeToTrackerTransform = 0.956528 -0.263235 0.125549 -190.803 0.268591 0.962861 -0.0275255 -97.8538 -0.113641 0.0600503 0.991706 -1948.58 0 0 0 1
Seq_Frame0001_ProbeToTrackerTransformStatus = OK
…
Seq_Frame0096_StylusToTrackerTransform = 1 0.000362791 -8.01739e-005 -0.106056 -0.000363818 1 -0.000369765 -0.839888 8.00139e-005 0.000370927 1 0.0862409 0 0 0 1
Seq_Frame0096_StylusToTrackerTransformStatus = INVALID
Seq_Frame0096_Timestamp = 355.783014
Seq_Frame0096_ImageStatus = OK
ElementDataFile = LOCAL
xœìÝw...

## Slide 9: Ultrasound image orientation
Specify the orientation of the X and Y axes of the image
Standard orientation in Plus: B-mode => MF, RF-mode => FM

## Slide 10: Temporal calibration
Goal: determine time offset between data streams
Extract the same information (e.g., position) and find the time offset that leads to maximum correlation of the signals
Use accurate clock (1ms is achievable with multimedia timers)
Timestamp filtering: detect delayed data, reduce jitter
Limit maximum speed
without temporal calibration
with temporal calibration
Frame index
time

## Slide 11: Spatial calibration
Goal: determine the PROBE to IMAGE transform

## Slide 12: Spatial calibration
Goal: determine the PROBE to IMAGE transform
Determine STYLUS to STYLUS TIP transform (pivot calibration)
Determine PHANTOM to PHANTOM REFERENCE transform (landmark registration)
Determine the IMAGE to PHANTOM transform (known fiducial line positions in the phantom, segmented visible line positions in the image)
PHANTOM
PHANTOM REFERENCE
STYLUS
STYLUS TIP
PROBE
IMAGE
TRACKER

## Slide 13: Spatial calibration
Tutorials with all data, models, tips&tricks.
Performing tracked ultrasound probe calibration using fCal
How to build an fCal calibration phantom

## Slide 14: Spatial calibration
Generic description of the phantom

## Slide 15: Processing
RF to B-mode conversion: brightness conversion and scan conversion for linear and curvilinear transducers
Volume reconstruction: with hole filling, multi-threaded
Ground truth
Reconstructed volume without hole filling
Reconstructed volume with hole filling

## Slide 16: fCal application
Spatial and temporal calibration
Basic visualization
Recording
Volume reconstruction

## Slide 17: PlusServer application
GUI and console application
Typically runs on the data collection computer
Broadcasts data through OpenIGTLink (3D Slicer, Matlab, etc.)
Remote control access
  Control data recording
  Request volume reconstruction
  Update transforms

## Slide 18: Performance
Calibration accuracy
  Temporal: standard deviation of repeated lag measurements 3 msec
  Spatial: fiducial line reprojection mean error 0.5mm (Ultrasonix L14-5 transducer with SonixGPS), total error 1mm
Data acquisition: fCal can typically keep up with the data acquisition speed (30-60 fps full-frame, 100-150 fps reduced depth and/or sector size)
OpenIGTLink broadcasting and visualization in 3D Slicer at 10-30 fps, total system latency 150ms for video input (latency of ultrasound imaging is typically an additional 80ms)

## Slide 19: Development process
Source control, tickets, releases, messaging (www.assembla.com/spaces/plus)
Standardized build environment (using CMake): automatically download and configure all required software components
Documentation: wiki, doxygen
Automatic tests: CTest, CDash, Sikuli

## Slide 20
System overview
Custom phantoms
hardware
Ultrasound scanner
OpenIGTLink
Navigation system
OpenIGTLink
3D Slicer extension manager
Registration
Segmentation
...
3D Slicer
Temporal and spatial calibration
Visualization
SlicerIGT
Transform recorder
Live ultrasound
PLUS
software
SlicerRT
DICOM-RT import
Dose comparison
VTK, ITK, CTK, QT, DCMTK, …
...
...
ProstateNav
Perk Tutor
MRI scanner
Custom needle guide
Developed in the PerkLab
Developed by collaborators
Device interface
Notes: what our research lab is doing from a software development perspective

## Slide 21
PLUS Architecture Overview
Hardware
Temporal calibration
Device
Algorithms
N-Wire calibration
Volume reconstruction
Single wall calibration
Line segmentation
Data
Tracked frame
Channel
Data sources
…
Device Specific SDK
Applications
fCal
PlusServer
iCal
…
Notes: This shows a layered approach to what depends on what
SDK – software development kit, functions to interact with a specific piece of hardware

## Slide 22
PLUS Data Pipeline
EM Tracker
Ultrasound
B-mode Channel
Channel
Rf-mode Channel
B Virtual Mixer
Rf Virtual Mixer
Rf Virtual Capture
B Virtual Capture
Data Flow
Channel
Channel
Visualization Client
Devices are the creation point of data
  Recorded data from world
  Simulated data, processed data
Channels are the communication point between devices
  Outputs tracked frames for client consumption

## Slide 23
PLUS: https://www.assembla.com/spaces/plus
Source code, binary releases, documentations, tutorials, etc.
Slicer IGT: https://www.assembla.com/spaces/SlicerIGT
3D Slicer extensions for additional visualization and analysis.
Contact: Andras Lasso (lasso@queensu.ca) http://perk.cs.queensu.ca
Thank you!

## Slide 24: Appendix

## Slide 25
Source: https://subversion.assembla.com/svn/plus/trunk/doc/overview/PlusOverview.pptx